cofrac

Accreditation N° 1-1114

Section Laboratoires

Laboratory Section

ANNEXE TECHNIQUE

TECHNICAL ANNEX

à l’attestation N° 1-1114 rev. 6

## to certificate No. 1-1114 rev. 6

L'accreditation concerne les prestations realisees par :

Certification concerns the services performed by:

### ACOME

Usine de Mortain/Romagny BP 45

### 50140 MORTAIN

Dans son unite technique :

In its technical unit:

### - LABORATOIRE D'ESSAIS

TEST LABORATORY

Elle est accordee selon le perimetre suivant :

It is granted according to the following test scope:

### ELECTRICITE/ Tout equipement ou produit electrique et/ou electronique soumis à des essais de comportement au feu

ELECTRICITY / All electrical and/or electronic equipment or products subject to fire behaviour testing

/ Essais de comportement au feu (77-2)

Fire behaviour tests (77-2)

Pour tous les essais concernant cette accreditation / For all tests concerning this accreditation :

(*) Le /aboratoire est reconnu competent pour pratiquer les essais en suivant les methodes referencees et leurs revisions ulterieures (FLEX1) / The laboratory is recognised as competent to carry out tests following the reference methods and their subsequent revisions (FLEX1).

La liste exhaustive des normes mises en œuvre est tenue à jour par le laboratoire / The exhaustive list of standards implemented is kept up to date by the laboratory.

Accreditation No. 1-1114

| Test object | Characteristics or quantities measured (including measuring range) | Test method principle (including main test means) | Test method reference (*) | Adaptable parameters (limits of adaptation of the method) |
| --- | --- | --- | --- | --- |
| Flame retardancy FE10 Conductors, electric or fiber optic cables | Burn height (0 to 60 cm) Flame height Temp. Propane flow rate Air flow rate Propane and air back pressure Time | Tests for classifying conductors and cables with regard to their fire behaviour - tests with the 1kW premix burner: measuring flame propagation along a sample of cable, placed vertically in a metal screen and subjected to the action of a 1kW flame for a predefined time | NFC 32070 test 1 C 32078-1 EN 60332-1-1 EN 60332-1-2 EN 60332-1-3 IEC 60332-1-1 IEC 60332-1-2 IEC 60332-1-3 + Other technical specifications (ORANGE / FT R&D) ACPE 75 method 8 | -Number of test tubes tested (origin and ageing) -Number of satisfactory test tubes |

ACPE: Acome Test Procedure

Accreditation No. 1-1114

| Test object | Characteristics or quantities measured (including measuring range) | Test method principle (including main test means) | Test method reference (*) | Adaptable parameters (limits of adaptation of the method) |
| --- | --- | --- | --- | --- |
| Fire retardancy FE110 Conductors, electric or fiber optic cables (in bundles) | Burn height (0 to 100 cm) Temperature Oven temperature Oven temp gradient Air speed in the oven Calibration curve in T0 Time | Test for classifying conductors and cables with regard to their fire behaviour – Ring furnace test: measuring flame propagation along a test tube (bundle of cable(s)) subjected to the heat of a ring furnace (+ pilot flames) in a ventilated cabinet, for a predefined time | NFC 32070 Test 2 + Other technical specifications (railway sector: SNCF, RATP) ACPE 135 | Burn height |
|  | Burn height (0 to 300 cm) Temperature Propane flow rate Burner air flow rate Ventilation air flow rate Time | Fire reaction test on cables arranged in layers: After a predefined application | C 32078-3 IEC 60332-3-10 IEC 60332-3-21 to 25 EN 60332-3-10 EN 60332-3-21 to 25 + Other technical specifications (railway sector: SNCF, RATP) ACPE 130 | Burn height |
| Fire retardancy |  | time, measuring flame propagation |  |  |
|  |  | along a layer of cables (predefined volume |  | Excepted class A |
| FE120 |  | of combustible materials) arranged |  | test with at least |
|  |  | vertically and subjected to the flame |  | one 35 mm2 |
| Conductors, electric or fiber optic cables (in layers) |  | of a 20.5 kW burner in a ventilated cabinet |  | conductor and requiring the use of |
|  |  |  |  | wide scale and |
|  |  |  |  | 2 burners |

ACPE: Acome Test Procedure

Accreditation No. 1-1114

| Test object | Characteristics or quantities measured (including measuring range) | Test method principle (including main test means) | Test method reference (*) | Adaptable parameters (limits of adaptation of the method) |
| --- | --- | --- | --- | --- |
|  | Same as above+ Heat output (0 to 500 kW) Calorific output THR (0 to 200 MJ) FIGRA (0 to 1500 W/s) Smoke opacity: SPR (0 to 2 m2/s) TSP (0 to 500 m2) Air flow Temperature Flaming droplets | Test for European classification of | EN 50399 2006/751/EC (European Commission decision of 27- 10-2006) ACPE 130 | Classification according to EN 13501-6 with additional classifications based on EN 61034 and EN 50267 Class A excepted |
|  |  | electrical conductors and cables with |  |  |
|  |  | regard to their fire behavior within |  |  |
| Fire retardancy |  | the framework of the Construction |  |  |
| FE120 |  | Products Regulation (CPR): After |  |  |
|  |  | a predefined burner application time, |  |  |
|  |  | determining flame propagation, |  |  |
| Conductors, electric or fiber optic cables |  | heat, calorific value and opacity of fumes emitted by a sheet of cables |  |  |
|  |  | (a layer of non-adjoining cables) |  |  |
|  |  | installed vertically on a metal |  |  |
|  |  | rack and exposed to a flame of |  |  |
|  |  | predefined strength (Class A excepted) |  |  |
| Smoke opacity FE70 Conductors, electric or fiber optic cables | Opacity (It: 0 to 100%) Temperature Volumes Air flow Time |  | C 32073-1 | Opacity and classification according to EN 13501-6 |
|  |  |  | C 32073-2 |  |
|  |  | Measurement of the density of fumes | EN 61034-1 |  |
|  |  | emitted by electric cables under defined conditions: Appling a defined heat source to one or more cables placed on a horizontal metal rack within a 27 m3 chamber, and measuring the optical decay after a defined period of time | EN 61034-2 IEC 61034-1 IEC 61034-2 + Other technical specifications (railway sector: |  |
|  |  |  | SNCF, RATP) |  |
|  |  |  | ACPE 250 |  |

ACPE: Acome Test Procedure

Accreditation No. 1-1114

| Test object | Characteristics or quantities measured (including measuring range) | Test method principle (including main test means) | Test method reference (*) | Adaptable parameters (limits of adaptation of the method) |
| --- | --- | --- | --- | --- |
| Smoke corrosivity FE80 / FE85 Insulating material or cable sheathing | pH (1 to 14) Conductivity (< 200 µSiem) Temperature Air flow Time | Conventional determination of smoke corrosivity: Determining corrosivity (measurement of pH and conductivity) of gases (trapped by impinging in purified water) emitted from a sample of pyrolysis material in a tube furnace at a predefined temperature | NFC 20453 C 32074 EN 50267-1 EN 50267-2-2 EN 50267-2-3 IEC 60754-2 ACPE 245 | Conductivity and classification according to EN 13501-6 |
|  |  | Conventional determination of the toxicity of a material from the | NF C20454 NF X 70100-1 | Qualitative research |
|  |  | quantity of toxic gases (CO, CO2, | NF X 70100-2 | with SEM and |
|  | CO emitted (0 to 500 mg) | halogenated acids, etc.) released | NF F 16101 | EDAX analyzer |
| Toxicity | CO2 emitted (0 to 2 g) | during its combustion: Based on | EN 50267-1 |  |
| FPF10 / FPF15 | HCI, HBr in solution (0 to 500 mg) | the pyrolysis of a material (insulation or | EN 50267-2-1 | CTI: Conventional |
|  | Temperature | cable sheathing) in a tubular furnace at a | IEC 60754-1 | Toxicity |
| Insulating material or cable | Air flow | predefined temperature, determining its | + | Index |
| sheathing |  | conventional toxicity index (CTI) from | Other specifications |  |
|  | Time | the quantities of CO et CO2 emitted, | PSA, RSA, ORANGE / FT | Analysis limited to CO, CO2, |
|  |  | and the halogens released | R&D and railway | HCI and HBr |
|  |  | (halogenated acids trapped in acid form |  |  |
|  |  | in an impinger solution) | ACPE 259 |  |

ACPE: Acome Test Procedure

# Accreditation made compulsory in the French regulatory framework specified by the text referenced in the Cofrac LAB INF 99 document available on www.cofrac.fr

Accreditation N° 1-1114

Date de prise d’effet / granting date : 07/02/2019	Date de fin de validité / expiry date : 31/05/2020

Cette annexe technique annule et remplace l'annexe technique 1-1114 Rev. 5.

This Technical Annex cancels and replaces Technical Annex 1-1114 Rev. 5